## Data Quality Documentation
Data Quality Statement
Analytical performance was deemed generally acceptable based on estimated surrogate recoveries and replicate precision. Out of a 130 analyte panel, 45% were present in the samples and fulfilled all QC requirements, whereas 30% of the analytes were not detected. 9-HODE and 13-KODE were above the highest calibration standard and results were estimated by linear extrapolation of the calibration curve and should therefore only be used qualitatively.
Table 1: Estimated surrogate recoveries |  | Table 2: Data quality status of measured metabolites (Note: Screens indicate pseudo-quantitative results)
Surrogate | Average Recoveries |  | Pass All QC |  | Background> 33% in 50-85% of samples |  | Observed above highest calibration |  | Non-Detects |  | Analytical Failure |  | Analytical replicate precision. Three samples were processed in duplicate. High abundance compounds generally showed high precision with ~60% of residues >100 pmol/g tissue having <35% RSD.
d4-6-Keto PGF1a | 34.9 ± 14% | 1 | 9,12,13-TriHOME | 1 | 6-keto-PGF1a | 1 | 9-HODE | 1 | RvE1 | 1 | PGE2 | Alkaline unstable
d4-TXB2 | 36.8 ± 12% | 2 | F2IsoP | 2 | 8,9-DiHETE | 2 | 13-KODE | 2 | TXB2 | 2 | PGD2 | Alkaline unstable
d4-PGF2a | 12 ± 3.8% | 3 | 15-keto-PGE2 | 3 | 14,15-DiHETrE |  |  | 3 | PGF2a | 3 | PGE2_EA | Alkaline unstable
d5-LXA4 | 55.4 ± 21% | 4 | RvD1 | 4 | 13,14-DiHDoPE |  |  | 4 | LXB4 | 4 | PGD2_EA | Alkaline unstable
d4-LTB4 | 57.5 ± 7.9% | 5 | 15,16-DiHODE_(screen) | 5 | 8,9-DiHETrE |  |  | 5 | RvD2 | 5 | Acetaminophen | Alkaline unstable
d11-14,15-DiHETrE | 63.8 ± 11% | 6 | 12,13-DiHODE_(screen) | 6 | 7,8-DiHDoPE |  |  | 6 | Naproxen | 6 | 1-AG | Alkaline unstable
d6-20-HETE | 48 ± 9.5% | 7 | 9,10-DiHODE_(screen) | 7 | 5,6-DiHETrE |  |  | 7 | LXA4 | 7 | 2-AG | Alkaline unstable
d4-9(S)-HODE | 15 ± 2.6% | 8 | 17,18-DiHETE | 8 | 15-KETE |  |  | 8 | RvE2 | 8 | 2-LG | Alkaline unstable
d3-13-KODE | 22.8 ± 4.9% | 9 | 14,15-DiHETE | 9 | 11(12)-EpETE |  |  | 9 | LTB5 | 9 | 1-LG | Alkaline unstable
d8-12-HETE | 33.2 ± 5.1% | 10 | 11,12-DiHETE | 10 | 12-HETE |  |  | 10 | 8,15-DiHETE | 10 | 2-OG | Alkaline unstable
d8-5(S)-HETE | 39.8 ± 7.7% | 11 | 12,13-DiHOME | 11 | 8-HETE |  |  | 11 | Protectin DX | 11 | 1-OG | Alkaline unstable
d4-12(13)-EpOME | 26.9 ± 4.7% | 12 | 9,10-DiHOME | 12 | 9-HETE |  |  | 12 | 5,15-DiHETE | 12 | 13-HpODE_(screen) | Retention time shift
d8-C20:4n6 | 26.2 ± 8.7% | 13 | 5,6-DiHETE | 13 | 5-KETE |  |  | 13 | Maresin_1
d4-PGF2a_EA | 124 ± 48% | 14 | 11,12-DiHETrE | 14 | N-C20:5n3_EA |  |  | 14 | 6-trans-LTB4
d8-C20:4n6_EA | 21.6 ± 7.1% | 15 | 12(13)-Ep-9-KODE | 15 | N-C20:4n6_Ser |  |  | 15 | LTB4
d8-C20:4n6_Gly | 23.9 ± 5.9% | 16 | 9-HOTE | 16 | N-C20:4n6_EA |  |  | 16 | Ibuprofen
d4-C16:0_EA | 16 ± 5.8% | 17 | 13-HOTE | 17 | N-C18:2n6_Gly |  |  | 17 | 19,20-DiHDoPE
|  | 18 | 18-HEPE | 18 | N-C22:4n6_EA |  |  | 18 | 16,17-DiHDoPE
|  | 19 | 9,10-DiHO | 19 | N-C18:1n9_Gly |  |  | 19 | 10,11-DiHDoPE
|  | 20 | 12-HEPE |  |  |  |  | 20 | 15-deoxy_PGJ2
|  | 21 | 9-HEPE |  |  |  |  | 21 | 15-HEPE
|  | 22 | 12-HODE |  |  |  |  | 22 | 20-HETE
|  | 23 | 10-HODE |  |  |  |  | 23 | 17-HDoHE
|  | 24 | 13-HODE |  |  |  |  | 24 | 14-HDoHE
|  | 25 | 5-HEPE |  |  |  |  | 25 | 12-KETE
|  | 26 | 9-HpODE_(screen) |  |  |  |  | 26 | 7-HDoHE
|  | 27 | 17(18)-EpETE |  |  |  |  | 27 | 4-HDoHE
|  | 28 | 15(16)-EpODE_(screen) |  |  |  |  | 28 | 8(9)-EpETrE
|  | 29 | 9(10)-EpODE_(screen) |  |  |  |  | 29 | 17-KDoHE
|  | 30 | 15-HETE |  |  |  |  | 30 | PGF2a_EA
|  | 31 | 12(13)-EpODE_(screen) |  |  |  |  | 31 | 19(20)-EpDoPE_EA
|  | 32 | 14(15)-EpETE |  |  |  |  | 32 | 11(12)-EpETrE_EA
|  | 33 | 9-KODE |  |  |  |  | 33 | N-C22:6n3_EA
|  | 34 | 11-HETE |  |  |  |  | 34 | N-C22:6n3_Gly
|  | 35 | 8(9)-EpETE |  |  |  |  | 35 | N-C20:4n6_Gly
|  | 36 | 19(20)-EpDoPE |  |  |  |  | 36 | N-C20:3n6_EA
|  | 37 | 5-HETE |  |  |  |  | 37 | N-C16:0_Gly
|  | 38 | 12(13)-EpOME |  |  |  |  | 38 | N-C20:0_EA
|  | 39 | 14(15)-EpETrE |  |  |  |  | 39 | N-C24:1n9_EA
|  | 40 | 9(10)-EpOME
|  | 41 | 16(17)-EpDoPE
|  | 42 | 13(14)-EpDoPE
|  | 43 | 10(11)-EpDoPE
|  | 44 | 11(12)-EpETrE
|  | 45 | 7(8)-EpDoPE
|  | 46 | 4-KDoHE
|  | 47 | 9(10)-EpO
|  | 48 | C20:5n3_(screen)
|  | 49 | C18:3n3_(screen)
|  | 50 | C22:6n3_(screen)
|  | 51 | C20:4n6_(screen)
|  | 52 | 18:2n6_(screen)
|  | 53 | N-C18:3n3_EA
|  | 54 | N-16:1n7_EA
|  | 55 | N-C18:2n6_EA
|  | 56 | N-C16:0_EA
|  | 57 | N-C18:1n9_EA
|  | 58 | N-C18:0_EA

## Acquisition Parameters
Table S1: Negative Mode Scheduled Multi Reaction Monitoring Experiment a | [7 empty] | Table S2: Positive Mode Scheduled Multi Reaction Monitoring Experiment a | [7 empty] | Table S3: API Scource Parameter Table
Compound | Acq Mid Point (min) | Q1 (Da) | Q3 (Da) | Window (sec) | Dwell Wt | DCP | CE | Compound | Acq Mid Point (min) | Q1 (Da) | Q3 (Da) | Window (sec) | Dwell Wt | DCP | CE | Parameter | (-)ESI Exp | (+)ESI Exp
PUHA_(-) | 1.84 | 249.2 | 130.1 | 12 | 0.5 | -30 | -18 | d4-Acetaminophen | 0.79 | 156.1 | 114.1 | 12 | 0.5 | 55 | 21 | CUR: | 20 | 20
RvE1 | 2.16 | 349.2 | 195.2 | 12 | 5 | -50 | -22 | Acetaminophen | 0.8 | 152.1 | 110.1 | 12 | 1 | 55 | 21 | IS: | -4500 | 4500
d4-6-Keto PGF1a | 2.18 | 373.3 | 199.2 | 12 | 0.5 | -70 | -20 | PUHA_(+) | 1.74 | 251.2 | 114.1 | 12 | 0.5 | 65 | 21 | TEM: | 650 | 650
6-keto-PGF1a | 2.18 | 369.3 | 163.1 | 12 | 2 | -70 | -33 | d4-PGF2a_EA | 2.18 | 384.3 | 62.1 | 12 | 0.5 | 50 | 42 | GS1: | 60 | 60
d4-TXB2 | 2.895 | 373.3 | 173.1 | 32 | 1 | -50 | -45 | PGF2a_EA | 2.19 | 380.3 | 62.1 | 12 | 2 | 45 | 39 | GS2: | 50 | 50
TXB2 | 2.9 | 369.3 | 169.1 | 32 | 2 | -50 | -21 | PGE2_EA + PGD2_EA | 2.2 | 378.3 | 62.1 | 18 | 2 | 65 | 40 | CAD: | Medium | Medium
9,12,13-TriHOME | 2.99 | 329.2 | 211.2 | 15 | 1 | -60 | -30 | CUDA_(+) | 5.7417 | 341.3 | 216.2 | 13 | 0.5 | 50 | 24 | EP | -10 | 10
d4-PGF2a | 3.03 | 357.2 | 197.2 | 12 | 0.5 | -70 | -33 | 19(20)-EpDoPE_EA | 7.6217 | 388.3 | 370.3 | 13 | 3 | 40 | 8.5 | CXP | -13 | 12
PGF2a | 3.05 | 353.2 | 193.2 | 30 | 3 | -70 | -33 | 11(12)-EpETrE_EA | 8.3017 | 364.3 | 62.1 | 13 | 3 | 75 | 45
PGE2 + PGD2 | 3.22 | 351.202 | 271.2 | 18 | 3 | -40 | -22 | N-C18:3n3_EA | 9.8817 | 322.2 | 62.1 | 13 | 1 | 60 | 21
LXB4 | 3.27 | 351.2 | 221.2 | 12 | 3 | -55 | -22 | N-C20:5n3_EA | 10.0017 | 346.2 | 62.1 | 13 | 1 | 70 | 39
d4-PGD2 | 3.23 | 355.2 | 275.2 | 12 | 0.5 | -40 | -27 | N-16:1n7_EA | 10.6217 | 298.2 | 62.1 | 13 | 1 | 105 | 21
RvD2 | 3.3 | 375.3 | 277.2 | 12 | 3 | -50 | -18 | N-C20:4n6_Ser | 11.4917 | 392.3 | 106.1 | 13 | 1 | 70 | 18
15-keto-PGE2 | 3.3 | 349.2 | 331.3 | 12 | 2 | -50 | -15 | N-C22:6n3_EA | 11.6217 | 372.3 | 62.1 | 13 | 1 | 55 | 45
d3-Naproxen | 3.4 | 232.1 | 188.1 | 12 | 0.5 | -35 | -9 | d8-C20:4n6_EA | 11.7117 | 356.3 | 63.1 | 13 | 0.5 | 50 | 45
Naproxen | 3.43 | 229.1 | 185.1 | 12 | 1 | -35 | -12 | N-C18:2n6_EA | 11.7617 | 324.2 | 62.1 | 13 | 1 | 70 | 21
RvD1 | 3.71 | 375.3 | 215.1 | 12 | 2 | -40 | -26 | N-C20:4n6_EA | 11.8017 | 348.2 | 62.1 | 13 | 1 | 70 | 39
d5-LXA4 | 3.74 | 356.5 | 222.1 | 12 | 0.5 | -40 | -20 | N-C22:6n3_Gly | 12.0017 | 386.3 | 311.2 | 13 | 1 | 50 | 15
LXA4 | 3.77 | 351.2 | 217.2 | 12 | 3 | -45 | -27 | d8-C20:4n6_Gly | 12.0817 | 370.3 | 76.1 | 13 | 0.5 | 45 | 21
RvE2 | 4.22 | 333.2 | 253.2 | 12 | 2 | -60 | -20 | N-C20:4n6_Gly | 12.1817 | 362.3 | 76.1 | 13 | 1 | 65 | 21
LTB5 | 4.68 | 333.2 | 195.2 | 12 | 2 | -50 | -21 | N-C18:2n6_Gly | 12.1817 | 338.3 | 76.1 | 13 | 1 | 90 | 17
15,16-DiHODE | 5.1 | 311.2 | 235.2 | 12 | 1 | -55 | -27 | N-C20:3n6_EA | 12.73 | 350.2 | 62.1 | 12 | 2 | 35 | 36
12,13-DiHODE | 5.15 | 311.2 | 183.2 | 12 | 1 | -60 | -27 | d5-2-AG | 12.7958 | 384.3 | 287.2 | 18.5 | 0.5 | 60 | 21
9,10-DiHODE | 5.18 | 311.2 | 201.2 | 12 | 1 | -65 | -27 | d4-C16:0_EA | 12.8 | 304.2 | 62.1 | 12 | 0.5 | 90 | 18
8,15-DiHETE | 5.2 | 335.2 | 235.2 | 12 | 2 | -65 | -21 | N-C16:0_EA | 12.82 | 300.2 | 62.1 | 12 | 1 | 70 | 18
Protectin DX | 5.41 | 359.2 | 153.1 | 12 | 2 | -40 | -22 | 1-_&_2-AG | 12.8458 | 379.302 | 287.2 | 18.5 | 2 | 95 | 21
17,18-DiHETE | 5.45 | 335.2 | 247.2 | 12 | 2 | -70 | -24 | 1-_&_2-LG | 12.9458 | 355.202 | 263.2 | 18.5 | 0.5 | 40 | 12
5,15-DiHETE | 5.49 | 335.2 | 173.1 | 12 | 2 | -40 | -18 | N-C16:0_Gly | 13.2017 | 314.2 | 76.1 | 13 | 1 | 50 | 15
Maresin_1 | 5.56 | 359.2 | 250.2 | 12 | 2 | -55 | -20 | N-C18:1n9_EA | 13.32 | 326.2 | 62.1 | 12 | 1 | 105 | 21
6-trans-LTB4 | 5.59 | 335.201 | 195.2 | 12 | 1 | -85 | -21 | N-C22:4n6_EA | 13.42 | 376.3 | 62.1 | 12 | 2 | 105 | 45
14,15-DiHETE | 5.73 | 335.2 | 207.2 | 12 | 2 | -50 | -22 | N-C18:1n9_Gly | 13.6717 | 340.2 | 76.2 | 13 | 1 | 60 | 21
d4-LTB4 | 5.82 | 339.2 | 197.2 | 12 | 1 | -50 | -24 | 1-_&_2-OG | 14.0458 | 357.202 | 265.2 | 18.5 | 0.5 | 55 | 15
LTB4 | 5.84 | 335.202 | 195.2 | 12 | 1 | -50 | -24 | N-C18:0_EA | 14.35 | 328.2 | 62.1 | 12 | 1 | 100 | 21
CUDA_(-) | 5.85 | 339.4 | 214.2 | 12 | 0.5 | -70 | -30 | N-C20:0_EA | 14.98 | 356.2 | 339.2 | 12 | 5 | 180 | 25
11,12-DiHETE | 5.9 | 335.2 | 167.1 | 12 | 2 | -40 | -22 | N-C24:1n9_EA | 15.44 | 410.3 | 62.1 | 12 | 5 | 180 | 41
d3-Ibuprofen | 5.95 | 208.2 | 164.1 | 13 | 0.5 | -50 | -9 | a - A medium collision gas pressure setting and default enterance and exit potentials were fixed and used for all analytes.
Ibuprofen | 5.97 | 205.1 | 161.1 | 13 | 1 | -50 | -9
12,13-DiHOME | 6.07 | 313.2 | 183.2 | 12 | 1 | -70 | -30
8,9-DiHETE | 6.2 | 335.2 | 185.1 | 12 | 2 | -30 | -22
9,10-DiHOME | 6.41 | 313.3 | 201.2 | 12 | 1 | -65 | -27
d11-14,15-DiHETrE | 6.56 | 348.4 | 207.2 | 12 | 0.5 | -65 | -24
14,15-DiHETrE | 6.64 | 337.2 | 207.2 | 12 | 2 | -60 | -24
19,20-DiHDoPE | 6.65 | 361.3 | 273.2 | 12 | 2 | -80 | -24
5,6-DiHETE | 6.77 | 335.2 | 145.1 | 12 | 2 | -30 | -22
16,17-DiHDoPE | 6.95 | 361.3 | 233.2 | 12 | 3 | -70 | -22
11,12-DiHETrE | 7.11 | 337.3 | 167.1 | 12 | 2 | -55 | -27
13,14-DiHDoPE | 7.12 | 361.3 | 193.1 | 12 | 3 | -80 | -23.5
12(13)-Ep-9-KODE | 7.285 | 309.2 | 291.2 | 12 | 1 | -90 | -21
9-HOTE | 7.34 | 293.2 | 171.1 | 12 | 1 | -60 | -21
10,11-DiHDoPE | 7.38 | 361.3 | 153.1 | 12 | 2 | -55 | -24
13-HOTE | 7.45 | 293.2 | 195.2 | 12 | 1 | -70 | -24
18-HEPE | 7.45 | 317.201 | 259.2 | 12.5 | 2 | -55 | -17
8,9-DiHETrE | 7.57 | 337.2 | 127.1 | 12.5 | 3 | -65 | -27
15-deoxy_PGJ2 | 7.61 | 315.2 | 271.2 | 12 | 3 | -50 | -21
9,10-DiHO | 7.74 | 315.2 | 297.2 | 12 | 1 | -110 | -30
15-HEPE | 7.82 | 317.2 | 219.2 | 12 | 1 | -55 | -18
d6-20-HETE | 7.83 | 325.3 | 281.2 | 12 | 0.5 | -80 | -21
20-HETE | 7.87 | 319.2 | 275.2 | 12 | 3 | -95 | -24
7,8-DiHDoPE | 7.9097 | 361.3 | 189.1 | 13 | 2 | -70 | -23
12-HEPE | 8.12 | 317.2 | 179.1 | 13 | 1 | -45 | -18
5,6-DiHETrE | 8.22 | 337.2 | 145.1 | 13 | 1 | -55 | -24
9-HEPE | 8.29 | 317.201 | 167.2 | 13 | 2 | -45 | -18
12-HODE + 10-HODE | 8.55 | 295.201 | 183.1 | 18.5 | 1 | -160 | -26
13-HODE | 8.56 | 295.201 | 195.2 | 15 | 1 | -90 | -24
5-HEPE | 8.62 | 317.2 | 115.1 | 13 | 2 | -40 | -21
d4-9(S)-HODE | 8.65 | 299.2 | 172.1 | 13 | 0.5 | -85 | -27
9-HODE | 8.7 | 295.201 | 171.1 | 13 | 1 | -70 | -24
13-HpODE + 9-HpODE | 8.94 | 311.2 | 293.2 | 18.5 | 1 | -55 | -15
17(18)-EpETE | 8.81 | 317.2 | 259.2 | 13 | 3 | -55 | -15
d3-13-KODE | 8.83 | 296.2 | 114.1 | 13 | 0.5 | -120 | -27.5
15(16)-EpODE + 9(10)-EpODE | 8.84 | 293.201 | 275.2 | 25 | 1 | -65 | -18
13-KODE | 8.87 | 293.2 | 179.1 | 13 | 1 | -80 | -27
15-HETE | 8.89 | 319.201 | 219.2 | 13 | 1 | -55 | -18
17-HDoHE | 9.01 | 343.201 | 281.2 | 13 | 2 | -55 | -18
12(13)-EpODE | 9.02 | 293.2 | 183.2 | 13 | 1 | -50 | -24
15-KETE | 9.13 | 317.2 | 273.2 | 12 | 2 | -60 | -18
14(15)-EpETE | 9.21 | 317.2 | 207.2 | 13 | 3 | -45 | -18
9-KODE | 9.24 | 293.2 | 185.2 | 13 | 1 | -100 | -27
11-HETE | 9.26 | 319.201 | 167.1 | 13 | 1 | -45 | -21
14-HDoHE | 9.27 | 343.202 | 281.2 | 13 | 2 | -60 | -18
11(12)-EpETE | 9.4 | 317.202 | 167.2 | 16 | 3 | -40 | -18
d8-12-HETE | 9.36 | 327.2 | 184.2 | 13 | 0.5 | -60 | -21
12-HETE | 9.4485 | 319.2 | 179.1 | 12 | 1 | -60 | -21
8(9)-EpETE | 9.48 | 317.2 | 155.1 | 12 | 3 | -40 | -16
8-HETE | 9.5854 | 319.201 | 155.1 | 12 | 1 | -45 | -21
12-KETE | 9.67 | 317.201 | 273.2 | 11 | 2 | -30 | -25
7-HDoHE | 9.69 | 343.2 | 281.2 | 12 | 2 | -80 | -17
9-HETE | 9.75 | 319.202 | 167.1 | 13 | 1 | -60 | -18
d8-5(S)-HETE | 10.01 | 327.2 | 116.1 | 13 | 0.5 | -70 | -18
19(20)-EpDoPE | 10.1 | 343.303 | 281.2 | 13 | 3 | -60 | -18
5-HETE | 10.11 | 319.2 | 115.1 | 13 | 1 | -50 | -18
d4-12(13)-EpOME | 10.11 | 299.2 | 198.1 | 13 | 0.5 | -90 | -24
12(13)-EpOME | 10.18 | 295.202 | 195.2 | 13 | 1 | -85 | -21
14(15)-EpETrE | 10.35 | 319.202 | 219.2 | 13 | 2 | -50 | -18
9(10)-EpOME | 10.44 | 295.202 | 171.1 | 13 | 1 | -75 | -21
4-HDoHE | 10.6 | 343.204 | 110.1 | 18 | 5 | -60 | -20
16(17)-EpDoPE | 10.55 | 343.2 | 274.15 | 13 | 3 | -45 | -15
13(14)-EpDoPE | 10.67 | 343.2 | 234.1 | 13 | 3 | -55 | -15
10(11)-EpDoPE | 10.82 | 343.2 | 153.1 | 13 | 3 | -60 | -16
5-KETE | 10.87 | 317.2 | 203.2 | 13 | 2 | -75 | -27
11(12)-EpETrE | 10.93 | 319.203 | 167.1 | 13 | 2 | -40 | -21
7(8)-EpDoPE | 11.11 | 343.2 | 189.1 | 13 | 3 | -40 | -15
8(9)-EpETrE | 11.15 | 319.202 | 155.1 | 13 | 3 | -40 | -18
17-KDoHE | 11.32 | 341.201 | 297.2 | 13 | 3 | -50 | -17
4-KDoHE | 11.73 | 341.202 | 297.2 | 13 | 3 | -40 | -17
9(10)-EpO | 12.1 | 297.2 | 279.2 | 13 | 1 | -105 | -27
C20:5n3 | 12.88 | 301.2 | 257.15 | 12 | 1 | -60 | -15
C18:3n3 | 13 | 277.2 | 259.15 | 12 | 1 | -115 | -24
C22:6n3 | 13.65 | 327.2 | 283.15 | 12 | 1 | -45 | -15
d8-C20:4n6 | 13.83 | 311.2 | 267.15 | 12 | 0.5 | -60 | -18
C20:4n6 | 13.88 | 303.2 | 259.15 | 12 | 1 | -40 | -18
18:2n6 | 14.02 | 279.2 | 261.15 | 12 | 1 | -185 | -38
a - A medium collision gas pressure setting and default enterance and exit potentials were fixed and used for all analytes.
`

## LC Gradient
Table S4: Chromatographic Parameters
Time | % Solv A | % Solv B
0.01 | 75 | 25
0.25 | 75 | 25
0.75 | 60 | 40
1.5 | 58 | 42
4.25 | 50 | 50
6.75 | 43 | 57
8 | 39 | 61
10.5 | 36 | 64
11 | 33 | 67
13 | 20 | 80
13.3 | 15 | 85
13.85 | 5 | 95
14 | 0 | 100
15.5 | 0 | 100
15.68 | 75 | 25
16.68 | 75 | 25
1) Solvent A: 0.1% Acetic Acid 2) Solvent B: 90% Acetonitrile / 10% isopropanol 3) flow rate = 0.5 mL/min 4) column: 2.1 X 150mm, 1.7 µm Waters Acquity C18 BEH 5) column temp: 60°C. 6) Max pressure: 15000 psi. Expected: 80000-12000psi. 7) Sampe Temp: 15 °C

## Quantitation Parameters
Table S5: Quantitative parameters for (+/-)-ESI/MS/MS lipid mediator + NSAID analysis
Compound | tR (min) |  | Concentration Range (nM) | ISTD a
d4-Acetaminophen | 0.89 |  | [144 - 144] | PUHA_(+)
Acetaminophen | 0.9 |  | [5.716 - 255200] | d4-Acetaminophen
PUHA_(-) | 1.84 |  | [101 - 101]
PUHA_(+) | 1.84 |  | [101 - 101]
RvE1 | 2.16 |  | [0.01248 - 891.6] | d4-PGF2a
d4-6-Keto PGF1a | 2.18 |  | [24 - 24] | PUHA_(-)
6-keto-PGF1a | 2.18 |  | [0.07557 - 5398] | d4-6-Keto PGF1a
d4-PGF2a_EA | 2.28 |  | [24.9 - 24.9] | PUHA_(-)
PGF2a_EA | 2.29 |  | [0.3522 - 15720] | d4-PGF2a_EA
PGE2_EA | 2.3 |  | [0.354 - 15800] | d4-PGF2a_EA
PGD2_EA | 2.45 |  | [0.354 - 15800] | d4-PGF2a_EA
d4-TXB2 | 2.895 |  | [53.4 - 53.4] | PUHA_(-)
TXB2 | 2.9 |  | [0.07557 - 5398] | d4-TXB2
9,12,13-TriHOME | 2.99 |  | [0.02648 - 1891] | d4-PGF2a
d4-PGF2a | 3.03 |  | [25.1 - 25.1] | PUHA_(-)
PGF2a | 3.05 |  | [0.07899 - 5642] | d4-PGF2a
F2-Isoprostanes | [2.6-3.5] |  | [0 - 0] | d4-PGF2a
PGE2 | 3.1 |  | [0.07943 - 5674] | d4-PGD2
LXB4 | 3.21 |  | [0.05902 - 4216] | d5-LXA4
d4-PGD2 | 3.24 |  | [25.2 - 25.2] | PUHA_(-)
RvD2 | 3.25 |  | [0.03718 - 1660] | d5-LXA4
PGD2 | 3.26 |  | [0.07943 - 5674] | d4-PGD2
15-keto-PGE2 | 3.26 |  | [0.1433 - 10240] | d4-PGD2
d3-Naproxen | 3.4 |  | [107 - 107] | PUHA_(-)
Naproxen | 3.43 |  | [5.471 - 244200] | d3-Naproxen
RvD1 | 3.71 |  | [0.02324 - 1660] | d5-LXA4
d5-LXA4 | 3.74 |  | [25.2 - 25.2] | PUHA_(-)
LXA4 | 3.77 |  | [0.03972 - 2837] | d5-LXA4
RvE2 | 4.22 |  | [0.02503 - 1788] | d4-PGF2a
LTB5 | 4.68 |  | [0.05495 - 3925] | d4-LTB4
15,16-DiHODE | 5.1 | b | [0 - 0] | d11-14,15-DiHETrE
12,13-DiHODE | 5.15 | b | [0 - 0] | d11-14,15-DiHETrE
9,10-DiHODE | 5.18 | b | [0 - 0] | d11-14,15-DiHETrE
8,15-DiHETE | 5.2 |  | [0.02488 - 1777] | d11-14,15-DiHETrE
Protectin DX | 5.41 |  | [0.02427 - 1734] | d5-LXA4
17,18-DiHETE | 5.45 |  | [0.02986 - 2133] | d11-14,15-DiHETrE
5,15-DiHETE | 5.49 |  | [0.02488 - 1777] | d11-14,15-DiHETrE
Maresin_1 | 5.56 |  | [0.02427 - 1734] | d5-LXA4
6-trans-LTB4 | 5.59 |  | [0.026 - 1857] | d4-LTB4
14,15-DiHETE | 5.73 |  | [0.02488 - 1777] | d11-14,15-DiHETrE
d4-LTB4 | 5.82 |  | [66.1 - 66.1] | CUDA_(-)
LTB4 | 5.84 |  | [0.026 - 1857] | d4-LTB4
CUDA_(-) | 5.85 |  | [99.9 - 99.9]
CUDA_(+) | 5.85 |  | [99.9 - 99.9]
11,12-DiHETE | 5.9 |  | [0.04977 - 3555] | d11-14,15-DiHETrE
d3-Ibuprofen | 5.95 |  | [95.6 - 95.6] | CUDA
Ibuprofen | 5.97 |  | [5.543 - 247500] | d3-Ibuprofen
12,13-DiHOME | 6.07 |  | [0.02782 - 1987] | d11-14,15-DiHETrE
8,9-DiHETE | 6.2 |  | [0.01991 - 1422] | d11-14,15-DiHETrE
9,10-DiHOME | 6.41 |  | [0.02671 - 1908] | d11-14,15-DiHETrE
d11-14,15-DiHETrE | 6.56 |  | [25.7 - 25.7] | CUDA_(-)
14,15-DiHETrE | 6.64 |  | [0.02585 - 1846] | d11-14,15-DiHETrE
19,20-DiHDoPE | 6.65 |  | [0.07724 - 3448] | d11-14,15-DiHETrE
5,6-DiHETE | 6.77 |  | [0.01991 - 1422] | d11-14,15-DiHETrE
16,17-DiHDoPE | 6.95 |  | [0.06758 - 3017] | d11-14,15-DiHETrE
11,12-DiHETrE | 7.11 |  | [0.02585 - 1846] | d11-14,15-DiHETrE
13,14-DiHDoPE | 7.12 |  | [0.0662 - 2956] | d11-14,15-DiHETrE
12(13)-Ep-9-KODE | 7.285 |  | [0.02708 - 1934] | d3-13-KODE
9-HOTE | 7.34 |  | [0.02377 - 1698] | d4-9(S)-HODE
10,11-DiHDoPE | 7.38 |  | [0.06694 - 2988] | d11-14,15-DiHETrE
13-HOTE | 7.45 |  | [0.02377 - 1698] | d4-9(S)-HODE
18-HEPE | 7.45 |  | [0.02629 - 1878] | d8-12-HETE
8,9-DiHETrE | 7.57 |  | [0.02585 - 1846] | d11-14,15-DiHETrE
15-deoxy_PGJ2 | 7.61 |  | [0.03097 - 2212] | d3-13-KODE
19(20)-EpDoPE_EA | 7.73 |  | [0.1445 - 6450] | d8-C20:4n6_EA
9,10-DiHO | 7.74 |  | [0.1587 - 11340] | d11-14,15-DiHETrE
15-HEPE | 7.82 |  | [0.05258 - 3756] | d8-12-HETE
d6-20-HETE | 7.83 |  | [24.5 - 24.5] | CUDA_(-)
20-HETE | 7.87 |  | [0.08736 - 3900] | d6-20-HETE
7,8-DiHDoPE | 7.9097 |  | [0.06319 - 2821] | d11-14,15-DiHETrE
12-HEPE | 8.12 |  | [0.02629 - 1878] | d8-12-HETE
5,6-DiHETrE | 8.22 |  | [0.02585 - 1846] | d11-14,15-DiHETrE
9-HEPE | 8.29 |  | [0.02629 - 1878] | d8-12-HETE
11(12)-EpETrE_EA | 8.41 |  | [0.1387 - 6190] | d8-C20:4n6_EA
13-HODE | 8.56 |  | [0.02683 - 1916] | d4-9(S)-HODE
12-HODE | 8.34 |  | [0.05902 - 4216] | d4-9(S)-HODE
10-HODE | 8.61 |  | [0.05902 - 4216] | d4-9(S)-HODE
5-HEPE | 8.62 |  | [0.02629 - 1878] | d8-5(S)-HETE
d4-9(S)-HODE | 8.65 |  | [26.6 - 26.6] | CUDA_(-)
9-HODE | 8.7 |  | [0.02778 - 1984] | d4-9(S)-HODE
13-HpODE | 8.86 | b | [0 - 0] | d4-9(S)-HODE
17(18)-EpETE | 8.81 |  | [0.02629 - 1878] | d4-12(13)-EpOME
15(16)-EpODE | 8.66 | b | [0 - 0] | d4-12(13)-EpOME
9(10)-EpODE | 8.8 |  | [0 - 0] | d4-12(13)-EpOME
d3-13-KODE | 8.83 |  | [26.9 - 26.9] | CUDA_(-)
13-KODE | 8.87 |  | [0.02675 - 1911] | d3-13-KODE
12(13)-EpODE | 9.02 | b | [0 - 0] | d4-12(13)-EpOME
15-HETE | 8.89 |  | [0.0546 - 3900] | d8-12-HETE
17-HDoHE | 9.01 |  | [0.02917 - 2083] | d8-12-HETE
15-KETE | 9.13 |  | [0.08791 - 3925] | d3-13-KODE
14(15)-EpETE | 9.21 |  | [0.02629 - 1878] | d4-12(13)-EpOME
9-KODE | 9.24 |  | [0.02675 - 1911] | d3-13-KODE
11-HETE | 9.26 |  | [0.0273 - 1950] | d8-12-HETE
14-HDoHE | 9.27 |  | [0.02917 - 2083] | d8-12-HETE
d8-12-HETE | 9.36 |  | [24.4 - 24.4] | CUDA_(-)
9-HpODE | 8.94 | b | [0 - 0] | d4-9(S)-HODE
11(12)-EpETE | 9.4 |  | [0.02629 - 1878] | d4-12(13)-EpOME
12-HETE | 9.4485 |  | [0.0273 - 1950] | d8-12-HETE
8(9)-EpETE | 9.48 |  | [0.02629 - 1878] | d4-12(13)-EpOME
8-HETE | 9.5854 |  | [0.0273 - 1950] | d8-12-HETE
12-KETE | 9.67 |  | [0.02716 - 1940] | d3-13-KODE
7-HDoHE | 9.69 |  | [0.08128 - 3628] | d8-12-HETE
9-HETE | 9.75 |  | [0.05495 - 3925] | d8-12-HETE
N-C18:3n3_EA | 9.99 |  | [0.1742 - 7776] | d8-C20:4n6_EA
d8-5(S)-HETE | 10.01 |  | [24.4 - 24.4] | CUDA_(-)
19(20)-EpDoPE | 10.1 |  | [0.08128 - 3628] | d4-12(13)-EpOME
5-HETE | 10.11 |  | [0.0273 - 1950] | d8-5(S)-HETE
d4-12(13)-EpOME | 10.11 |  | [133 - 133] | CUDA_(-)
N-C20:5n3_EA | 10.11 |  | [0.1621 - 7236] | d8-C20:4n6_EA
12(13)-EpOME | 10.18 |  | [0.02683 - 1916] | d4-12(13)-EpOME
14(15)-EpETrE | 10.35 |  | [0.0546 - 3900] | d4-12(13)-EpOME
9(10)-EpOME | 10.44 |  | [0.02587 - 1848] | d4-12(13)-EpOME
4-HDoHE | 10.6 |  | [0.02917 - 2083] | d8-5(S)-HETE
16(17)-EpDoPE | 10.55 |  | [0.04064 - 1814] | d4-12(13)-EpOME
13(14)-EpDoPE | 10.67 |  | [0.07315 - 3266] | d4-12(13)-EpOME
N-16:1n7_EA | 10.73 |  | [0.1882 - 8403] | d8-C20:4n6_EA
10(11)-EpDoPE | 10.82 |  | [0.07315 - 3266] | d4-12(13)-EpOME
5-KETE | 10.87 |  | [0.08547 - 3816] | d3-13-KODE
11(12)-EpETrE | 10.93 |  | [0.0273 - 1950] | d4-12(13)-EpOME
7(8)-EpDoPE | 11.11 |  | [0.07315 - 3266] | d4-12(13)-EpOME
8(9)-EpETrE | 11.15 |  | [0.0546 - 3900] | d4-12(13)-EpOME
17-KDoHE | 11.32 |  | [0.07299 - 3259] | d3-13-KODE
N-C20:4n6_Ser | 11.6 |  | [0.143 - 6386] | d8-C20:4n6_Gly
4-KDoHE | 11.73 |  | [0.07358 - 3285] | d3-13-KODE
N-C22:6n3_EA | 11.73 |  | [0.1507 - 6728] | d8-C20:4n6_EA
d8-C20:4n6_EA | 11.82 |  | [28.1 - 28.1] | CUDA_(+)
N-C18:2n6_EA | 11.87 |  | [0.1731 - 7728] | d8-C20:4n6_EA
N-C20:4n6_EA | 11.91 |  | [0.1612 - 7194] | d8-C20:4n6_EA
9(10)-EpO | 12.1 |  | [0.02805 - 2004] | d4-12(13)-EpOME
N-C22:6n3_Gly | 12.11 |  | [0.1453 - 6485] | d8-C20:4n6_Gly
d8-C20:4n6_Gly | 12.19 |  | [27.1 - 27.1] | CUDA_(+)
N-C20:4n6_Gly | 12.29 |  | [0.1549 - 6916] | d8-C20:4n6_Gly
N-C18:2n6_Gly | 12.29 |  | [0.1659 - 7407] | d8-C20:4n6_Gly
N-C20:3n6_EA | 12.83 |  | [0.1602 - 7151] | d4-C16:0_EA
C20:5n3 | 12.88 | b | [0 - 0] | d8-C20:4n6
C18:3n3 | 13 | b | [0 - 0] | d8-C20:4n6
d5-2-AG | 12.95 |  | [26.1 - 26.1] | CUDA_(+)
d4-C16:0_EA | 12.9 |  | [26.4 - 26.4] | CUDA_(+)
N-C16:0_EA | 12.92 |  | [0.187 - 8347] | d4-C16:0_EA
2-AG | 12.82 |  | [0.2958 - 13210] | d5-2-AG
1-AG | 13.04 |  | [0.2958 - 13210] | d5-2-AG
2-LG | 12.92 |  | [0.3949 - 17630] | d5-2-AG
1-LG | 13.18 |  | [0.3949 - 17630] | d5-2-AG
N-C16:0_Gly | 13.31 |  | [0.1786 - 7974] | d8-C20:4n6_Gly
N-C18:1n9_EA | 13.42 |  | [0.172 - 7680] | d4-C16:0_EA
N-C22:4n6_EA | 13.52 |  | [0.1864 - 8320] | d4-C16:0_EA
C22:6n3 | 13.65 | b | [0 - 0] | d8-C20:4n6
N-C18:1n9_Gly | 13.78 |  | [0.1649 - 7364] | d8-C20:4n6_Gly
d8-C20:4n6 | 13.83 |  | [25.6 - 25.6] | CUDA_(-)
C20:4n6 | 13.88 | b | [0 - 0] | d8-C20:4n6
18:2n6 | 14.02 | b | [0 - 0] | d8-C20:4n6
2-OG | 14.08 |  | [0.4712 - 21040] | d5-2-AG
1-OG | 14.25 |  | [0.5278 - 23560] | d5-2-AG
N-C18:0_EA | 14.45 |  | [0.1709 - 7631] | d4-C16:0_EA
N-C20:0_EA | 15.08 |  | [0.1312 - 5859] | d4-C16:0_EA
N-C24:1n9_EA | 15.54 |  | [0.1553 - 6934] | d4-C16:0_EA
a - ISTD (i.e. interna standard) associations with analytes used to correct quantification for analytical loss and suppression b - analyte levels estimated by ISTD adjusted peak area, but without a quantitative response factor calibration